# SPeech Recognition In Noise Test (SPRINT) for H-3 Profile

# Mary T. Cord, Brian E. Walden and Rodney M. Atack

# Walter Reed Army Medical Center

# BACKGROUND

Despite the Army's extensive hearing conservation program, noise-induced hearing loss remains a significant health hazard to the soldier.  Currently, more than fifty thousand active duty Army personnel are rated permanent H-3 under the Army's physical profiling system.  Soldiers with H-3 profiles may be required to go before a Military Medical Retention Board (MMRB) which makes a determination to retain, reclassify, or separate the soldier from active duty, ostensibly on the basis of the soldier's ability to perform his/her duties.  However, it is clear that soldiers with H-3 profiles vary considerably in their ability to perform their mission. Further, MMRB's vary considerably in their decisions when assessing soldiers with H-3 profiles.  As a result, there has long been concern among Army audiologists, as well as local commanders, that some soldiers have been unnecessarily reclassified or separated from the Army, while others have been retained in assignments which may further jeopardize the soldier's fitness for duty and, therefore, mission accomplishment.

Although many soldiers with H-3 profiles experience little difficulty in favorable communication situations, there is great inter-subject variability in the amount of difficulty experienced by these soldiers when communicating in noisy listening environments such as in combat training or on the battlefield.  The degree of communication difficulty experienced by an individual soldier cannot be predicted with accuracy from traditional clinical audiometric measures such as pure-tone thresholds or speech recognition in quiet.  A clinical test that evaluates speech recognition ability in noise should have greater validity predicting communication handicaps that would impinge on mission accomplishment than the traditional audiometric indices that are currently provided to the MMRB.

A simple clinical test of speech recognition in noise has been developed and standardized which can be easily implemented by Army audiologists to provide MMRB's with information regarding a hearing-impaired soldier's potential communication handicap on the job.  The information provided should allow the Board to make more informed and equitable decisions regarding an individual soldier's fitness for duty.

# TEST DEVELOPMENT

The SPeech Recognition in Noise Test (SPRINT) consists of 200 monosyllabic words (Form C of the NU-6 lists) pre-recorded in a background of multi-talker babble noise.  The speech-to-babble ratio (+9 dB) is such that normal-hearing soldiers are little affected by the background noise; that is, they can identify 95-100% of the monosyllabic words correctly.  However, the performance of soldiers with H-3 profiles varies considerably.  Administration of the SPRINT allows the performance of an individual soldier to be compared with that of a large random sample of soldiers with H-3 profiles.

During the development of the SPRINT, a total of 319 soldiers with H-3 profiles were tested at 12 different Army audiology clinics within CONUS.  The data obtained were sent to WRAMC to be compiled and analyzed. Initially, the data from the 12 test sites were compared to determine if systematic differences existed among the test facilities.  An analysis of variance indicated no significant differences.  Consequently, data were pooled for all subsequent analyses.  The mean score for the 319 soldiers was 163 (out of a possible 200) with a range of 65 to 196 words correct.  A frequency distribution of the scores (i.e., number of words correct) for the 319 H-3 profile soldiers is shown in Figure 1.

# TEST ADMINISTRATION

The SPRINT should be administered by an audiologist or by a technician under an audiologist's supervision.

# Who to test:

The SPRINT is to be administered to all active duty soldiers who are seen through your clinic and are identified as having a permanent H-3 profile.  Additionally, all H-3 soldiers are to be re-tested with the SPRINT whenever a significant permanent threshold shift has occurred.  According to Army hearing conservation guidelines, this is defined as a decrease of more than 10 dB for the pure-tone average of 2000, 3000, 4000 Hz or a decrease of 15dB or more at 1000, 2000, 3000 or 4000 Hz, in either ear.

# Test Procedures:

The SPRINT is administered through an audiometer under earphones with the signal (i.e., monosyllabic words presented in a multi-talker babble) delivered to the right and left ear simultaneously.  Because the monosyllabic words and the background babble are pre-recorded at the proper speech-to-babble ratio on the same channel of the tape, the audiologist need only set the overall presentation level of the recording.

- Set audiometer attenuator to 50 dB HL.

- Route signal to both right and left earphones.

- Start tape and adjust VU to 0 (zero) for calibration tone.

- Turn down tape monitor and turn up talk-back to hear patient responses more clearly; position talk-back microphone close enough to patient that responses can be clearly heard.

# Instructions to soldier:

"You will be hearing a tape recording of a man's voice saying some words.  In the background, there will be the sound of several people talking at once.  I want you to repeat back the words that the man is saying.  Some of the words may be difficult for you to hear.  If you're not sure what the word is, take a guess.  There are 4 lists containing 50 words each for a total of 200 words.  This will take about 20 minutes.  You will be hearing the words and the background talk in both earphones."

- Be sure to encourage guessing.

- If more than one answer is given (such as "It was either map or mat") take the first response.

- Write down the incorrect responses on the response form.

- The complete 200 word test must be administered for this test to be valid.  Interpolating a total score from the first 50 or 100 words invalidates the test.

# SCORING

A soldier's SPRINT score is simply the number of monosyllabic words that were correctly identified (i.e., 0-200).  In order for an individual soldier's score to be ranked in comparison with the standardization sample, it is necessary to convert the soldier's score to a percentile ranking within the frequency distribution depicted in Figure 1.  This is accomplished using the table given in Appendix 1.  For example, if an individual soldier correctly identifies 133 of the 200 monosyllabic words, s/he falls in the 10th percentile and, therefore, 90% of all soldiers with H-3 profiles in the Army will score higher than s/he on the SPRINT.  Similarly, if a soldier got 182 of the 200 words correct, s/he falls in the 83rd percentile.  In this case, only 17% of all soldiers with H-3 profiles in the Army will score higher.

# RECOMMENDATIONS TO MMRB

The percentile score derived from the soldier's raw score indicates how he/she compares with all other soldiers in the Army with H-3 profiles.  Although this is useful information, it is not sufficient to make a recommendation to the MMRB.  Many other factors might be considered (e.g., age, progression of the hearing loss, military rank, duty assignment, etc.).  Typically, however, the most important consideration in addition to the SPRINT score will be the soldier's length-in-service.  For example, a soldier who falls in the 50th percentile who has just completed basic training must be considered differently from a soldier with the same percentile ranking who has 18 years on active duty.  At a minimum, therefore, the recommendation to the MMRB will be based on a combination of the soldier's percentile ranking and his/her length of service on active duty.  These two factors are combined to make a single recommendation, using the table depicted in Appendix 2.  This table defines five recommendation categories.  To determine the appropriate recommendation for an individual soldier, enter his/her percentile score along the horizontal axis.  Next, enter his/her length in service (in years) along the vertical axis.  The intersection of these two points within the table indicates the appropriate recommendation.  For example, if a soldier scored at the 55th percentile on the SPRINT and has 11.5 years of active duty service, s/he would fall within the "C" category.  A recommendation of "Re-assignment to non-noise hazardous duty," therefore, would be made.  As is apparent from the table in Appendix 2, soldiers who perform better on the SPRINT and who have more years on active duty are more likely to be recommended for retention than are soldiers who perform poorly on the SPRINT and have less time in service.

The audiologist's recommendation is to be recorded in the soldier's permanent medical records on a Standard Form 600 which should be attached to the most recent Physical Profile (DA Form 3349).  The suggested format for an overprint can be seen in Appendix 3.  Your command may require approval of overprints; check with your medical records department for guidance.  The SPRINT response form should be filed in the soldier's sub-specialty chart in the Audiology clinic.  The SPRINT response form is not a DA Form and therefore should not be filed in the soldier's permanent medical record.

# RELIABILITY

The reliability of the recommendations that result from using the SPRINT and the table in Appendix 2 was assessed by obtaining test-retest scores on the SPRINT from 30 soldiers.  The mean score for the initial administration of the SPRINT was 164.6 words correct (S.D. = 15.46), and the mean score for the retest was 166.8 words correct (S.D. = 14.21).  A product-moment correlation coefficient of .93 (p< .01) indicates a strong positive association between the test and retest scores. As was apparent from the mean SPRINT scores for the two administrations seen in Figure 2, there was a tendency for scores to improve on retest.  This was likely due to learning effects since, in most cases, retest occurred within 4 days of the initial test.  Although additional improvement in SPRINT scores may have resulted from further administrations of the test, it must be kept in mind that the sampling distribution upon which the percentile ranks are computed is based on a single administration of the SPRINT to the standardization sample.  There is no apparent reason to assume that the relative ranking of an individual soldier in comparison to all other soldiers with H-3 profiles would change with repeated administrations of the test.

Appendix 1	PERCENTILES for SPRINT

SCORE      PERCENTILE

0-75	1

76-100	2

101-107	3

108-116	4

117-122	5

123-124	6

125-126	7

127-128	8

129-130	9

131-135	10

136-138	11

139-140	12

141	13

142	14

143-144	16

145	17

146	18

147	19

148	21

149-150	22

151-152	23

153	24

154	26

155-156	28

157	29

158	30

159	32

160	34

161	35

SCORE       PERCENTILE

162	37

163	40

164	42

165	45

166	48

167	49

168	51

169	53

170	55

171	57

172	59

173	62

174	64

175	67

176	70

177	72

178	74

179	76

180	78

181	80

182	83

183	86

184	89

185	92

186-187	95

188	96

189	97

190	98

191	99

Appendix 2

CATEGORY	RECOMMENDATION

A	Retention in current assignment

B	Retention in current assignment with restrictions

C	Re-assignment to (or retention in) non-noise hazardous 	AOC/MOS

D	Discretionary **

E	Separation from service

**  For soldiers falling in category D, the audiologist can make a recommendation associated with any category adjacent to Category D. Except for patients with 18+ years on active duty (for which a Category B recommendation could be made), this choice will be between Category C (re-assignment) or Category E (separation).  The decision of which recommendation to make should be based on such factors as stability of loss, potential for further noise exposure, the soldier's AOC/MOS, and the recommendation of the local commander.

# SPRINT (SPeech Recognition In Noise Test) for soldiers with H-3 Profile

Name:

Date:

SSN:                                                     Rank:

Date of Birth:

Total number correct out of 200:                          (raw score)

Percentile:

(see Appendix 1 to convert raw score to %ile score)

Length in Service:                        yrs.                            mos.

# Category for Recommendation: (see Appendix 2)

# A - Retain in current assignment

# B - Retain in current assignment with restrictions

C - Re-assign to (or retain in) non-noise hazardous                            AOC/MOS

# D - Discretionary*

# E - Separation from service

* Discretionary - may recommend any adjacent category: C (re-assignment) or E (separation); or for those with 18+ years on active duty, B (retention with restrictions).

# Comments:

Audiologist:

# LIST 1C

1. raise

2. door

3. tip

4. sure

5. hurl

6. met

7. burn

8. sell

9. reach

10. dime

11. jar

12. death

13. which

14. third

15. pool

16. moon

17. fat

18. king

19. chalk

20. yes

21. week

22. whip

23. bean

24. choice

25. rag

26. fall

27. vine

28. jail

29. home

30. boat

31. mode

32. tough

33. lot

34. raid

35. take

36. page

37. keen

38. laud

39. limb

40. goose

41. gap

42. sub

43. nag

44. size

45. hash

46. love

47. knock

48. puff

49. shout

50. kite

# LIST 2C

1. dead

2. juice

3. merge

4. young

5. calm

6. bite

7. rain

8. match

9. book

10. loaf

11. nice

12. bought

13. ton

14. shawl

15. white

16. hate

17. shack

18. pike

19. fail

20. rot

21. gin

22. pad

23. gaze

24. live

25. room

26. south

27. mill

28. which

29. tool

30. numb

31. haze

32. pick

33. turn

34. goal

35. voice

36. keep

37. thought

38. far

39. read

40. hush

41. chair

42. chief

43. keg

44. soap

45. said

46. dab

47. wag

48. deep

49. learn

50. lore

# LIST 3C

1. mop

2. tell

3. germ

4. seize

5. good

6. base

7. search

8. ring

9. half

10. mess

11. late

12. when

13. gun

14. name

15. pain

16. ditch

17. rat

18. hire

19. shall

20. cheek

21. pole

22. life

23. date

24. road

25. talk

26. cause

27. youth

28. bar

29. lid

30. hit

31. cool

32. five

33. rush

34. phone

35. mouse

36. thin

37. pearl

38. beg

39. walk

40. luck

41. team

42. soup

43. dodge

44. chat

45. sheep

46. note

47. void

48. jug

49. wire

50. cab

# LIST 4C

1. perch

2. bath

3. back

4. bone

5. wife

6. fit

7. shirt

8. wash

9. neat

10. tire

11. make

12. mob

13. doll

14. pass

15. sour

16. dog

17. vote

18. time

19. near

20. lease

21. rose

22. kill

23. came

24. food

25. lose

26. should

27. check

28. kick

29. long

30. tape

31. yearn

32. lean

33. gas

34. sail

35. red

36. wheat

37. have

38. mood

39. dip

40. such

41. hall

42. rough

43. peg

44. get

45. chain

46. join

47. ripe

48. judge

49. hole

50. thumb